## ChangeLog
Version | Date | Author | Comments
0.1 | 13.10.2011 | PF | creation, draft version for internal review
1 | 27.10.2011 | PF | Issue for delivery to STScI

## During cooldown-stabilization
Date | Start time | End time | Activities / Envents | Status | Comments | CSV file(s) | Range of exposure IDs
ACTIVITIES CONDUCTED DURING COOLDOWN AND THE TEMEPRATURE STABILIZATION
30.01.2011 | 22:00 | 4:52 | FILL-DET-DARK-SHORT #01 6 series of 5 "short" dark exposures (88 groups). This test sequence is a "filler" for unused night time during the final days of the bench coold down. As such it can be (and was) interrupted whenever needed. | Interrupted after 25 darks | Procedure executed while the bench was stabilizing. No quick-look performed on-site. Script ran OK but CAA keywords were set to 'UNKNOWN' (the ICE was probably OFF). | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/FILL-DET-DARK-SHORT-01/FILL-DET-DARK-SHORT.csv | Day2011030/4091-4102 Day2011031/4103-4115
31.01.2011 | 23:30 | 8:15 | FILL-DET-DARK-SHORT #02 6 series of 5 "short" dark exposures (88 groups). This test sequence is a "filler" for unused night time during the final days of the bench coold down. As such it can be (and was) interrupted whenever needed. | Interrupted after 27 darks | Procedure executed while the bench was stabilizing. No quick-look performed on-site. Script ran OK but there was a problem with the SITS IGSS. Last exposure was ingested later than the others. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/FILL-DET-DARK-SHORT-02/FILL-DET-DARK-SHORT.csv | Day2011031/4116-4117 Day2011032/4118-4142
01.02.2011 | 23:30 | 8:12 | FILL-DET-DARK-SHORT #03 6 series of 5 "short" dark exposures (88 groups). This test sequence is a "filler" for unused night time during the final days of the bench coold down. As such it can be (and was) interrupted whenever needed. | Interrupted after 28 darks | Procedure executed while the bench was stabilizing. No quick-look performed on-site. Script ran OK but CAA keywords were set to 'UNKNOWN' (the ICE was probably OFF). | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/FILL-DET-DARK-SHORT-03/FILL-DET-DARK-SHORT.csv | Day2011032/4143 Day2011033/4144-4170
02.02.2011 | 23:17 | 8:30 | FILL-DET-DARK-SHORT #04 6 series of 5 "short" dark exposures (88 groups). This test sequence is a "filler" for unused night time during the final days of the bench coold down. As such it can be (and was) interrupted whenever needed. | Interrupted after 11 darks | Procedure executed while the bench was stabilizing. No quick-look performed on-site. Procedure was aborted due to a SITS crash during exposure 4185. The final SWTS cleanup created exposure 4186. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/FILL-DET-DARK-SHORT-04/FILL-DET-DARK-SHORT.csv | Day2011033/4174-4176 Day2011034/4177-4184

## Preparation phase
Date | Start time | End time | Activities / Envents | Status | Comments | CSV file(s) | Range of exposure IDs
ACTIVITIES CONDUCTED DURING THE PREPARATION PHASE
03.02.2011 | 15:00 |  | Actual start of the preparation phase. - instrument and test equipment ready - Detector temperature control ON with a set point of 38.5 - Bench at 34 K (corresponding to an interface temperature of 31 K) | n/a |  |  | n/a
03.02.2011 | 15:10 | 16:15 | PREP-INIT Measure count rate in specific configurations in order to determine the number of groups for preparation phase exposures. | OK | None. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-INIT-01/PREP-INIT.csv | Day2011034/4187-4203
03.02.2011 | 16:15 | 17:40 | PREP-CLS-FSM-A Derive the nominal position of the FSM. | NOT OK | Script ran flawlessly but examination of the exposures showed that the field stop mask (FSM) did not come back to its original position after a large amplitude deplacement along the x-axis (-10 mm). Displacement along the x and y-axes are observed. The (0,0) position of the FSM is not the nominal position anymore. An NCR has been issued. Important note: as a consequence the FSM positions that are found in the FITS headers cannot be trusted. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-CLS-FSM-A-01/PREP-CLS-FSM-A.csv | Day2011034/4204-4219
03.02.2011 | 17:40 | 19:30 | PREP-DEBUG-FSM Trouble-shooting procedure aiming at understanding what had happened to the CMO/FSM during the PREP-CLS-FSM-A test sequence. | Trouble-shooting | Trouble-shooting exposures aiming at deriving the new nominal position. Discussion with ASD experts: we hit a hard stop; cryo-mechanism (CMO) is fully functional but zero position is lost. Derived the new zero position (-1.400mm;-1.207mm). CAUTION: the OBS-ID used in the procedure (PREP-DEBUG-CMO) differs from the procedure name. | Not available. | Day2011034/4220-4223
03.02.2011 | 19:35 | 20:00 | PREP-CLS-FSM-B - RUN #01 Check that the FSM zero position (derived from the data acquired during part-A) is correct. | OK | Checked that the new zero position is correct => OK. CAUTION: update of the CMO database not possible before 04.02.2011 in the afternoon; In the mean-time this offset must be applied in all procedures. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-CLS-FSM-B-01/PREP-CLS-FSM-B.csv | Day2011034/4224
03.02.2011 | 20:10 | 21:31 | PREP-CLS-PHM-A A blind dither pattern is applied trying to locate the pinhole in the fixed-slits and in the IFU field of view. Part A of the full CLS/PHM setup procedure. | OK | Script ran flawlessly with blind dither around the (corrected) nominal position to try to locate the pinhole in the A1600 aperture and in the IFU field of view. The groups of 3 pinholes designed for the IFU are clearly seen in the IFU field of view. The pair of pinholes designed for the A1600 location are clearly seen in one of exposures. The determination of the nominal pinhole mask (PHM) on-going in parallel to the next activities. CAUTION: CMO database not yet updated and an offset is applied in the procedure to center correctly the CMO. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-CLS-PHM-01/PREP-CLS-PHM-A.csv | Day2011034/4235-4251
03.02.2011 | 21:40 | 22:40 | PREP-RCSS-A A blind dither pattern is applied trying to locate the RCSS point-source in the A1600 fixed-slit. Part A of the full RCSS setup procedure. | Failed to locate the RCSS source | Script ran flawlessly but it was not possible to locate the RCSS (single point source). We need to change strategy (new procedure will be written tomorrow). CAUTION: CMO database not yet updated and an offset is applied in the procedure to center correctly the CMO. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-RCSS-01/PREP-RCSS-A.csv | Day2011034/4235-4251
03.02.2011 | 23:00 | 23:45 | PREP-BACKGROUND-CHK Procedure to check the vacuum chamber background levels and the inpact of background radiation on the exposures (both imaging and spectrscopic modes). | OK | Script ran flawlessly; background rates compiled; As during the DM campaign, the background count rates prevent us from getting unsaturated full-frame images if we use the CLEAR aperture or one of the long-pass-filter. We have acceptable rates with the band-pass filters. This should not be a big issue for the calibration. It is confirmed that we need to plan for ON/OFF exposures for "red" spectroscopic configurations (those pushing up to 5 microns) but no show stopper. This stategy is already implemented in our procedures. CAUTION: CMO database not yet updated and an offset is applied in the procedure to center correctly the CMO. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-BKGD-CHECK-01/PREP-BKGD-CHECK.csv | Day2011034/4252-4275
04.02.2011 | 0:55 | 2:50 | MON-DET-DARK-SHORT Sequence of 10 short dark exposures (88 groups). | Procedure interrupted after 4 darks. | Some problems encountered during the script execution. The exposures look fine but we expect to have some persistence in some of them. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/MON-DET-DARK-SHORT-01/MON-DET-DARK-SHORT.csv | Day2011035/4276-4280
04.02.2011 | 3:43 | 11:04 | PREP-MSA-CONTRAST Measure the MSA contrast. | OK although the CAA lamp is off in some exposures. | We hit the maximum on-time (20 min) for the CAA lamps; quick-look indicates that the contrast images are probably already deep enough in a single MULTIACCUM 88x1 exposure => this should not be an issue, this is not a problem for "regular" use of the CAA lamaps (i.e. for everything but the MSA contrast exposures) It was decided not to interrupt the procedure (we will use the exposures with the lamp OFF for some quick-and-dirty persistence measurements). Note that we had some kind of a traffic jam between the SITS and it took ages to get the exposures on the BigMac1 | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-MSA-CONTRAST-01/PREP-MSA-CONTRAST.csv | Day2011035/4281-4300
04.02.2011 | 11:30 | 16:30 | PREP-CAA (trouble shooting: CAA light levels too high) This test aims at obtaining exposures with all CAA lamps using the appropriate medium resolution gratings, the PRISM, and the MIRROR, in order to obtain reliable count-rates for these lamps and their dependence on drive current. This data is needed to derive the needed attenuation factors for the CAA lamps, as it is currently too bright. Important note: this test sequence included the use of side-B electronics for parts E and F. | OK trouble-shooting | Got data on the count rates for all the CAA modes at 4 different current levels and for the primary side Got data on the count rates for all CAA modes at 2 different current levels and for the redundant side. Preliminary analysis points toward typical differences of 10% between the primary and redundant sources with a maximum of 25%. This is good news. Inputs for the preparation of the fixes that will be implemented between the two cryogenic cycles. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-CAA-01/PREP-CAA-A.csv | Part A: Day2011035/4301-4315 Part B: Day2011035/4316-4330 Part C: Day2011035/4331-4345 Part D: Day2011035/4346-4360 Part E (side B): Day2011035/4361-4375 Part F (side B): Day2011035/4376-4390
CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-CAA-01/PREP-CAA-B.csv
CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-CAA-01/PREP-CAA-C.csv
CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-CAA-01/PREP-CAA-D.csv
CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-CAA-01/PREP-CAA-E.csv
CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-CAA-01/PREP-CAA-F.csv
04.02.2011 | 17:00 | 20:15 | PREP-CLS-XCAL (CLS trouble-shooting) This test aims at getting the CLS FFV and FFB modes into flux ranges usable for MOS/S200 slits by changing the position of FW3 in the CLS lightbox. | OK trouble-shooting | Fixed the issue using a different attenuation factor. Got the cross-calibration data. Starting after this trouble-shooting sequence the default configuration for the FFV and FSB has changed. Note: part A has been run twice. CAUTION: CMO database not yet updated and an offset is applied in the procedure to center correctly the CMO. CAUTION: baseline FW3 position for the FFV (from 7 to 4) and FFB (from 3 to 1) has changed on the 04.02.2011 20:15. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-CLS-XCAL-01/PREP-CLS-XCAL-A.csv | Part A (run #01): Day2011035/4391 Part A (run #02): Day2011035/4392 Part B: Day2011035/4393-4398 Part C: Day2011035/4399-4405 Part D: Day2011035/4406-4413 Part E: Day2011035/4414-4420
CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-CLS-XCAL-02/PREP-CLS-XCAL-A.csv
CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-CLS-XCAL-B-01/PREP-CLS-XCAL-B.csv
CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-CLS-XCAL-C-01/PREP-CLS-XCAL-C.csv
CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-CLS-XCAL-D-01/PREP-CLS-XCAL-D.csv
CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-CLS-XCAL-E-01/PREP-CLS-XCAL-E.csv
04.02.2011 | 20:15 | 20:15 | Update of the CMO database Update performed to bring back the database to its original state before the failure that took place during the PREP-CLS-FSM-A execution (03/02/2011 16:15). | n/a | n/a
04.02.2011 | 20:15 | 20:30 | PREP-CLS-FSM-B (RUN #02) Check of the positioning of the CMO/FSM. | OK trouble-shooting | Checked that the update of the CMO database performed during the afternoon brought us back where we were yesterday at the beginning of the preparation phase. The image of the field stop in the MSA plane is back to within 1/2 micro-shutter width (1 detector pixel) => OK Important note: after that the CMO position indications have been restored to their original values. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-CLS-FSM-B-01/PREP-CLS-FSM-B.csv | Day2011035/4421
04.02.2011 | 19:30 | 21:15 | PREP-CLS-PHM-B This test aims at checking the CMO/PHM coordinates that position pinholes in the A1600 aperture and the IFU field of view. | OK | Script interrupted after the exposure at nominal PHM position due to an error of the CMO; procedure restarted OK. Control images with the nominal PHM position derived from part A (-0.68mm;+0.85mm) that should put two pinholes at the center of A1600 and 9 pinholes within the IFU field of view) Perfect, we even see the pinhole in the other (much narrower) fixed-slits. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-CLS-FSM-B-02/PREP-CLS-FSM-B.csv | Day2011035/4422-4425
04.02.2011 | 21:15 | 22:45 | NASA/ADS trouble shooting activities | n/a |  |  | n/a
04.02.2011 | 22:45 | 23:35 | PREP-CLS-PHM-C (RUN #01) This test aims at checking the readout parameters (number of groups) for the PHM-based configurations that will be used during the calibration campaign. | NOT OK | Script executed by error (only PREP-CLS-PHM-B should have been executed) and the PHM positions had not been updated. Script to be updated and the procedure run again. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-CLS-PHM-C/PREP-CLS-PHM-C.csv | Day2011035/4426-4431
04.02.2011 | 23:50 | 2:15 | PREP-CLS-ULS This test aims at deriving the readout parameters (number of groups) for the ULS-based configurations that will be used during the calibration campaign. | OK | Script executed flawlessly; nice sharp single emission line at 2.8 microns as expected. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-CLS-ULS-01/PREP-CLS-ULS.csv | Day2011036/4432-4441
05.02.2011 | 2:30 | 4:15 | PREP-FOCUS This test aims at deriving the NIRSpec best focus position for the nominal OTE focus position early on in the NIRSpec test campaign. | OK | Script executed flawlessly. Nice focus curves on several pinholes yielding a nominal focus position of -280 microns. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-FOCUS-01/PREP-FOCUS.csv | Day2011036/4442-4460
05.02.2011 | 4:20 | 6:00 | MON-MSA-FPA-RAD #01 Monitoring exposures (limited scope for cycle1): failed-open map of the MSA; detector cross-hatching. | NOT OK | Script executed flawlessly There was an error in the requested configuration and the exposures were obtained with the pinhole mask in place and not the field stop mask; exposures are not usable. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/MON-MSA-FPA-RAD-01/MON-MSA-FPA-RAD.csv | Day2011036/4461-4472
05.02.2011 | 6:15 | 10:55 | PREP-CONFOCALITY-CHK This test verifies that all NIRSpec observing modes share the same focus at cryogenic temperatures. | OK | Script executed flawlessly; focus sweeps were performed over the -500,+500 regions without recentering on the best focus position (-280 micron), we should change that in the next version of the procedure. Focus curves have been generated and are all consistent with the value derived with the F140X filter. Very little sensitivity to focus in the red. Note: the used OBS-ID prefix (PREP-FOCCHK) differs from the procedure name. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-CONFOCALITY-CHK-01/PREP-CONFOCALITY-CHK.csv | Day2011036/4473-4535
05.02.2011 | 11:24 | 15:54 | PREP-DET-CHK-1 This test sequences will allow probing the bias level, dark level, the noise level and the behavior of the reference pixels. This is a persistence sensitive test. | OK | Script executed with some problems (without any impact on the data itself): - it took ages before the exposure reach the BigMac1 (getting a lot of successive 88-group exposures really stresses the system); -the "superbias" exposure uncovered a bug in the script that caused it to stop (but the exposure was obtained) Preliminary analysis yields the expected typical dark current levels. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-DET_CHECK-1-01/PREP-DET-CHECK-1.csv | Day2011036/4536-4546
05.02.2011 | 16:02 | 18:15 | PREP-DET-CHK-2 This second test sequence aims at checking the level of cross-hatching that is present on the two SCAs. | OK | Script executed with some problems on the DHAS. The ingestion was stuck and was successfully restarted. Once more, the system does not really like long series of 8-group exposures. There should be no impact on the data. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-DET_CHECK-2-01/PREP-DET-CHECK-2.csv | Day2011036/4547-4551
05.02.2011 | 18:30 | 20:00 | PREP-DET-CHK-3 This third test sequence aims at probing the gain value. | OK | Script executed witout any problem. Count rates are in the right range for gain checks. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-DET_CHECK-3-01/PREP-DET-CHECK-3.csv | Day2011036/4552-4571
05.02.2011 | 20:00 | 21:54 | PREP-DET-CHK-4 This fourth test sequence aims at verifying that there are no glitches in the histogram (as seen during the detector system level testing), verifying the saturation level and if the linearity parameters derived during at DS level are still applicable. | OK | Script executed witout any problem. Count rates are in the right range for linearity and saturation level checks as well as for checking that no histogram "glitch" is present. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-DET_CHECK-4-01/PREP-DET-CHECK-4.csv | Day2011036/4572-4576
05.02.2011 | 22:00 | 22:45 | ASD work on the CMO to increase the accessible range along the z-axis - range increased from 1.5mm to 2.0mm | n/a |  |  | n/a
05.02.2011 | 22:53 | 0:20 | PREP-RCSS-SEARCH Trouble-shooting activities to try to find the RCSS source after failing to do so during PREP-RCSS-A (03.02.2011 21:40). | Failed to locate the RCSS source | RCSS source was seen but subsequent attempts to put it in the A1600 aperture failed. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-RCSS-SEARCH-01/PREP-RCSS-SEARCH.csv | PREP-RCSS-SEARCH Day2011036/4577-4597
06.02.2011 | 0:20 | 0:45 | PRE-PRCSS-WIN (RUN #01) Trouble-shooting activities to try to find the RCSS source after failing to do so during PREP-RCSS-A (03.02.2011 21:40) and PREP-RCSS-SEARCH (06.02.2011 22:53). New procedure inserted according to PVS #56. | Failed to locate the RCSS source | The procedure was aborted after a single exposure as it was obvious that it was not correct. We need time to review the procedure and to understand what went wrong. Starting a different procedure in the mean time. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-RCSS-WIN-01/PREP-RCSS-WIN.csv | PREP-RCSS-WIN (run #1): Day2011036/4598
06.02.2011 | 0:50 | 1:40 | PREP-CLS-PHM-C This test sequence aims at obtaining imaging exposures for the CMO/PHM configurations in order to derive the corresponding readout parameters. | OK | Script ran flawlessly; pinholes are present => OK. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-CLS-PHM-C/PREP-CLS-PHM-C.csv | Day2011037/4599-4604
07.02.2011 | 1:50 | 2:15 | PRE-PRCSS-WIN (RUN #02) Trouble-shooting activities to try to find the RCSS source after failing to do so during PREP-RCSS-A (03.02.2011 21:40) and PREP-RCSS-SEARCH (06.02.2011 22:53). New procedure inserted according to PVS #56. Run #02 as the previous one (07.02.2011 0:20) was aborted. | Failed to locate the RCSS source | Once more, we did not manage to position the RCSS in the A1600 aperture. We will have to try again tomorrow. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-RCSS-WIN-02/PREP-RCSS-WIN.csv | Day2011037/4605-4609
06.02.2011 | 2:15 | 13:00 | IFU-COMBO1 (CALIBRATION PHASE ACTIVITY) - PVS #57 This test aims at obtaining all exposures necessary to perform the distortion, flat-field, radiometric and wavelength calibration of the spectrographic stage of the IFU for all gratings and the PRISM. In addition the data will allow us to calibrate the CAA (continuum and spectral line sources) and to verify the Erbium reference source when used in IFU mode. | Problem with the Argon lamp for some configurations | Script ran without problem. A problem with the Argon lamp detected (in 50 % of the exposures the Argon lamp is almost inexistant), NCR raised. Some configurations will have to be repeated during the calibration phase. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/IFU-COMBO1-01/IFU-COMBO1.csv | Day2011037/4610-4759
06.02.2011 | 13:15 | 22:10 | IMA-FOCUS (CALIBRATION PHASE ACTIVITY) - PVS #57 This test verifies that NIRSpec can be focused with the RMA over the entire OTE focus range. | OK | There was a problem during the execution of the script; error corrected and script restarted (hence the very long effective duration) - nominal focus position for z=0 mm: -295 microns (with a plateau covering roughly 100 microns) - nominal focus position for z=2 mm: +247 microns (with a plateau covering roughly 100 microns) - nominal focus position for z=-3.1 mm: around -1100 microns (no fit available yet) | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/IMA-FOCUS-01/IMA-FOCUS.csv | Day2011037/4760-4816
06.02.2011 | 22:20 | 23:005 | PREP-RCSS-A (RUN #02) A blind dither pattern is applied trying to locate the RCSS point-source in the A1600 fixed-slit. Part A of the full RCSS setup procedure. | OK | Another attempt to position the RCSS in the A1600 slit after a series of failures. Script ran smoothly and the RCSS was present in the central image; nominal RCSS coordinates determined (18.00,-16.54) | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-RCSS-A-01/PREP-RCSS-A.csv | Day2011037/4817-4828
07.02.2011 | 23:27 | 23:41 | PREP-RCSS-B Check that the nominal RCSS source position derived from the part-A data is correct. | OK | Position is OK. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-RCSS-B-01/PREP-RCSS-B.csv | Day2011037/4829
06.02.2011 | 0:00 | 0:30 | PREP-FOCUS-RCSS This test aims at trying to find the location of the RCSS through the MSS by using one of the brightest RCSS sources in PRISM or imaging modes. | OK | We realized that the RCSS may not be confocal with the FSM and PHM. A focus sweep was performed, yielding a nominal focus position around -700 microns (with a large plateau). Value as been implemented in part C of PREP-RCSS. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-FOCUS-RCSS-01/PREP-FOCUS-RCSS.csv | Day2011037/4830-4846
06.02.2011 | 1:00 | 2:00 | PREP-RCSS-C This test sequence consists in MOS/SLIT spectrographic exposures with the RCSS positioned in the A1600 aperture for all the RCSS configurations in order to derive the corresponding readout parameters. | OK | Script ran smoothly; no surprise here; list of peak count rates available in the log. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-RCSS-C-01/PREP-RCSS-C.csv | Day2011038/4847-4868
07.02.2011 | 2:00 | 11:40 | PREP-NINT-TEST This test aims to assess whether taking an exposure with multiple integrations is equivalent to taking multiple exposures with a single integration. | OK | Script ran smoothly; we were afraid of running into trouble with the data flow between the SITS and the BigMac1 but this went well. There was a quite large wait-time for the data (~1h at max) but the DHAS did not get stalled Only quick look performed, analysis will be performed off-line. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-NINT-TEST-01/PREP-NINT-TEST.csv | Day2011038/4869-4892
07.02.2011 | 11:50 | 13:30 | MON-MSA-FPA-RAD (RUN #02) Monitoring exposures (limited scope for cycle1): failed-open map of the MSA; detector cross-hatching. | OK | This second was run a second time after correcting the inputs to the script to make sure we have the CMO field-stop in position and to avoid the problem encountered during run #01. Only quick-look performed, analysis will be performed off-line. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/MON-MSA-FPA-RAD-02/MON-MSA-FPA-RAD.csv | Day2011038/4893-4904
07.02.2011 | 13:30 | 15:45 | In parallel: EGSE troubleshooting and test of the Argon source by ASD | n/a |  |  | n/a
07.02.2011 | 15:45 | 16:50 | PREP-CLS-ULS This test aims at deriving the readout parameters (number of groups) for the ULS-based configurations that will be used during the calibration campaign. | OK | Was forgotten in the original timeline report. Nothing reported. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-CLS-ULS-01/PREP-CLS-ULS.csv | Day2011038/4905-4915
07.02.2011 | 16:50 | 17:50 | PREP-MSA-HEATUP-A (control exposures before heat-up) Acquisition of the imaging mode control exposures. We use the CAA/TEST source as well as a much brighter source to use the contrast map to monitor for internal distortion and potentially contrast changes. Acquisition of the spectrographic mode control exposures using external sources as the CAA sources are currently too bright. | OK | Sfter updating the inputs to the script, we obtained the imaging mode and G140H control exposures BEFORE the heat-up Quick-look performed and data is OK. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-MSA-HEAT-UP-A-01/PREP-MSA-HEAT-UP-A.csv | Day2011038/4916-4921
07.02.2011 | 17:50 | 21:15 | MSA HEAT-UP ITSELF Execution of the actual heat-up procedure using a script form Astrium. A total of 6 exposure are obtained as part of this script (see PREP-MSA-HEAT-UP.doc) for a description of these files. | OK | Quite impressive heat-up with the 260K temperature reached for all quadrants in less than 13min; MSA quadrants were held to > 260K during 15min before completely stopping the heater and starting the cool-down Cool-down from 260K to 60K in 4 hours A set of 6 exposures taken by the script itself and stored in PREP-MSA-HEAT-UP-PROC. |  | Day2011038/4922-4927 (NRS_MSA-HU-pre0-XX)
07.02.2011 | 21:15 | 22:03 | PREP-MSA-HEATUP-B (control exposures after heat-up) Acquisition of the post heat-up spectrographic mode control exposures. | OK | We obtained the G140H and imaging mode exposures AFTER the heat-up Quick comparison of the data before and after the heat-up was performed; one failed-open has now closed and we also see that some low contrast regions with partially closed micro-shutters have seen their contrast slightly improved (no quantitative values available yet) Overall, the heat-up did improve things - no obvious shift detected when comparing the images before and after (i.e. this would confirm the good stability of the system) - analysis to be performed off-line | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-MSA-HEAT-UP-B-01/PREP-MSA-HEAT-UP-B.csv | Day2011038/4928-4932
07.02.2011 08.02.2011 | 22:04 | to well within the next day | PREP-MSA-HEATUP-C (control exposures during the end of the cool-down) Monitoring of the positioning and internal geometry of the MSA while its temperature drops down from 60 K to its initial operating temperature. | OK | Series of control exposures during the end of the cool-down Only quick look performed for the first runs, analysis to be performed off-line. Starting mid-day of the 08.02.2011, no quick look performed (no more ESA scienc people on site). CAUTION: one more run performed between the EMC-EMI periods (together with the PREP-MSA-HEAT-UP-D part. See next tab. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-MSA-HEAT-UP-C-01/PREP-MSA-HEAT-UP-C.csv | #01 Day2011038/4933-4934 #02 Day2011038/4935-4936 #03 Day2011039/4937-4938 #04 Day2011039/4939-4940 #05 Day2011039/4941-4942 #06 Day2011039/4943-4944 #07 Day2011039/4945-4946 #08 Day2011039/4947-4948 #09 Day2011039/4949-4950 #10 Day2011039/4951-4952 #11 Day2011039/4953-4954
08.02.2011 | during day | during day | MSA inspection High-resolution pictures of the MSA through the cross-piece with front illumination. | n/a | Conducted by J.-F. Pittet in the morning. Exact time missing but must have been inserted between two runs of the PREP-MSA-HEATUP-C procedure. |  | n/a

## EMC-EMI phase
Date | Start time | End time | Activities / Envents | Status | Comments | CSV file(s) | Range of exposure IDs
ACTIVITIES CONDUCTED BETWEEN THE PREPARATION AND CALIBRATION PHASES (EMC/EMI NIGHTS MAINLY)
08.02.2011 | 20:21 | 5:35 | FILL-DET-PERSISTENCE-1 (A) Persistence test conducted in imaging mode with the FLAT1 source (we are not allowed to turn the GWA before the last MSA cool-down check has been performed). We will have saturation in 1 frames in the slits and saturation in 10-15 frames in low contrast regions. | OK | Script ran without problem. No on-site quick-look performed. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/FILL-DET-PERSISTENCE-A-01/FILL-DET-PERSISTENCE_1.csv | Day2011039/4955-4968 Day2011040/4969-4983
09.02.2011 | 6:30 | 7:00 | PREP-MSA-HEATUP-C (control exposures during the end of the cool-down) Monitoring of the positioning and internal geometry of the MSA while its temperature drops down from 60 K to its initial operating temperature. | OK | No problem reported. No on-site quick-look performed. CAUTION: the other runs have been performed at the end of the preparation phase. See previous tab. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-MSA-HEAT-UP-C-01/PREP-MSA-HEAT-UP-C.csv | Day2011040/4984-4985
09.02.2011 | daytime | daytime | EMC/EMI testing | n/a | Some [dark] exposures taken (no description available here). |  | Day2011040/4986-4988 (NRS_DS_CE_)
09.02.2011 | 6:30 | 7:00 | PREP-MSA-HEATUP-D Final set of control exposures after the MSA heat-up. | OK | No problem reported. No on-site quick-look performed. CAUTION: the heat-up procedure was executed earlier on during the preparation phase. See previous tab. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/PREP-MSA-HEAT-UP-D-01/PREP-MSA-HEAT-UP-D.csv | Day2011040/4989-4990
09.02.2011 | 20:54 | 5:37 | FILL-DET-PERSISTENCE-2 (B) Persistence test using the CLS/FF2 +G235H configuration and aiming at reaching saturation slowly (88 frames). | OK | Script ran without problem. No on-site quick-look performed. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/FILL-DET-PERSISTENCE-B-01/FILL-DET-PERSISTENCE_2.csv | Day2011040/4991-5001 Day2011041/5002-5019
10.02.2011 | daytime | daytime | EMC/EMI testing | n/a | Some [dark] exposures taken (no description available here). |  | Day2011041/5020-5069
10.02.2011 | 20:40 | 4:54 | FILL-DET-PERSISTENCE-3 (C) Persistence test using the CAA/FLAT2 +G235H configuration and aiming at quickly to reaching saturation (about 10-15 frames) and then keeping illuminating the detector for a total of 88 frames to match the exposure time in the other tests. | OK | Script ran without problem. No on-site quick-look performed. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/FILL-DET-PERSISTENCE-C-01/FILL-DET-PERSISTENCE_3.csv | Day2011041/5070-5081 Day2011042/5082-5098
11.02.2011 | daytime | daytime | EMC/EMI testing | n/a | Some [dark] exposures taken (no description available here). |  | Day2011042/5099-5153
11.02.2011 | 22:39 | 4:21 | DET-LIN This test aims to characterize the linearity and saturation level. | OK | Script ran without problem. No on-site quick-look performed. Data not yet analysed. | CYCLE1-DHAS-TREE/2_TVac1/2_Cal_Prep/DET_LIN-01/DET-LIN.csv | Day2011042/5154-5158 Day2011043/5159-5173
12.02.2011 | daytime | daytime | EMC/EMI testing | n/a | Some [dark] exposures taken (no description available here). |  | Day2011043/5174-5185

## Calibration phase
Date | Start time | End time | Activities / Envents | Status | Comments | CSV file(s) | Range of exposure IDs
ACTIVITIES CONDUCTED DURING THE CALIBRATION PHASE
20.02.2011 | 10:52 | 12:58 | MON-MSA-FPA-RAD #01 Monitoring exposures (limited scope for cycle1): failed-open map of the MSA; detector cross-hatching. | OK | Script ran without a problem. Quick-look performed on the 21/02/2011. Exposures are OK (including GWA-related keywords). |  | Day2011051/5190-5201
20.02.2011 | 12:21 | 15:20 | MON-DET-DARK-SHORT #01 Sequence of 10 short dark exposures (88 groups). | OK | Script ran without a problem. Quick-look performed on the 21/02/2011. |  | Day2011051/5202-5206
20.02.2011 | 15:26 | 17:24 | MON-DET-DARK-SHORT #02 Sequence of 10 short dark exposures (88 groups). | OK | Script ran without a problem. Ingestion and transfer to the BigMac1 was very very slow. Quick-look performed on the 21/02/2011. |  | Day2011051/5207-5211
20.02.2011 | 17:36 | 19:20 | CHECK-PHM-IFU Procedure to check that the CMO position placing the IFU pinhole in the middle of the IFU field of view is correct. | OK | The script was interrupted because of a problem with the CMO (NCR raised). The CMO was move "manually" into PHM position. Ingestion and transfer to the BigMac1 was very very slow till Xavier reduced the number of NFS threads to one on ther DHAS. Quick-look performed on the 21/02/2011. IFU image reconstruction shows that we are centered in the IFU field of view to within 1 slice. Note: ULS2 source was on during this procedure (no impact as we are in imaging mode with a much brighter broad-band source). | CYCLE1-DHAS-TREE/2_TVac1/5_Cal_Phase/CHECK-PHM-IFU-01/CHECK-PHM-IFU.csv | Day2011051/5212-5215
20.02.2011 | 19:20 | 21:30 | CMO DEBUGGING + BACK TO FSM POSITION | OK |  |  | n/a
20.02.2011 | 21:38 | 23:48 | IFU-COMBO2 Characterization of the spectrograph PSF at 2.8-microns in IFU mode using the CLS/ULS2 monochromatic source. | OK | Script ran without problem. No problem detected during the quick-look. |  | Day2011051/5216-5227
20.02.2011 | 23:57 | 1:53 | MON-DET-DARK-SHORT #03 Sequence of 10 short dark exposures (88 groups). | OK | Script ran without a problem. Quick-look performed on the 21/02/2011. |  | Day2011051/5228-5231 Day2011052/5232
21.02.2011 | 1:58 | 11:35 | DET-WINDOW-TEST Comparison of the performances of the detector when using the full-frame / window / stripe modes (hopefully, no difference). | OK | Script running without a problem however the 2048x2048 window-mode exposure has been automatically transformed into a full-frame exposure by the flight-software (NCR #1187 raised). Transfer between the DHAS and BigMac1 is once more very very slow. Note: this test sequence was rerun between the calibration and confirmation phases (test sequence name: DET-WINDOW-TEST-SUB). | CYCLE1-DHAS-TREE/2_TVac1/5_Cal_Phase/DET-WINDOW-TEST-01/DET-WINDOW-TEST.csv | Day2011052/5233-5323
21.02.2011 | 11:36 | 17:51 | PHM-PS-XCAL - Parts A to D (PVS #82) Before/after control exposures for the change of PSB, PS1, PS2 and PS3 baseline attenuation (going for lower attenuations and therefore higher electron rates). | OK | The four scripts ran without a problem. The baseline CLS lightbox wheel 3 positions for the PSB, PS1, PS2 and PS3 configurations were updated (now position 7 instead of 3). The gain in intensity was lower than expected (between 2 and 4 instead of the canonical 6.35) probably due to the apodisation of the optical beam in the lightbox. Important note: this change cannot be traced using the keywords except looking at the positions of the individual wheels (the CLS mode name remained the same). It can therefore easily go unnoticed. |  | Part A: Day2011052/5324-5343 Part B: Day2011052/5344-5361 Part C: Day2011052/5362-5379 Part D: Day2011052/5380-5399 CAUTION: Part-D exposures have an incorrect OBS-ID: PREP-CLS-XCAL-D (CLS instead of PS)
21.02.2011 | 18:11 | 19:53 | DET-QE-CUTOFF Characterization of the detector QE cutoff at long wavelengths. | OK | Script ran without a problem. No issue detected during the quicklook. | CYCLE1-DHAS-TREE/2_TVac1/5_Cal_Phase/DET-QE-CUTOFF-01/DET-QE-CUTOFF.csv | Day2011052/5400-5408
21.02.2011 | 19:56 | 21:29 | COMBO3 Characterization of the spectrograph PSF at 2.8-microns in MOS/SLIT mode using the CLS/ULS2 monochromatic source. | OK | Script ran without a problem. No issue detected during the quicklook. | CYCLE1-DHAS-TREE/2_TVac1/5_Cal_Phase/COMBO3-01/COMBO3.csv | Day2011052/5409-5418
21.02.2011 | 21:49 | 6:18 | SCI-OBS-SIM Simulation of observation at low intensity levels in IFU mode using the pinhole mask (continuum spectrum + ULS2 line). | OK | Script ran without a problem. No issue detected during the quicklook. |  | Day2011052/5419-5433 Day2011053/5434-5453
22.02.2011 | 6:30 | 8:00 | IMA-PSF RUN #01 Characterization of the imaging-mode PSF using the RCSS point source in the A1600 aperture. Measurements for various focus positions. | NOT OK | The selected RCSS source is much much too bright and the peak of the PSF saturates immediately. The script is stopped manually at 8:00. We will have to choose another source and rerun the script. Note: this script was rerun on the 23.02.2011. Note: ULS2 lamp in ON during this test. This is not an issue as it does not contribute when using the RCSS (different light source and the CLS optical path is blocked). |  | Day2011053/5454-5488
22.02.2011 | 8:20 | 14:10 | GWA-SENSOR-CALIB Characterization of the repeatability of the GWA for its different positions. Calibration of the GWA tilt sensors. | OK | Script ran without a problem (actually significantly faster than estimated). No issue detected during the quicklook. | CYCLE1-DHAS-TREE/2_TVac1/5_Cal_Phase/GWA-SENSOR-CALIB-01/GWA-SENSOR-CALIB.csv | Day2011053/5489-5568
22.02.2011 | 14:24 | 16:44 | MON-MSA-FPA-RAD #02 Monitoring exposures (limited scope for cycle1): failed-open map of the MSA; detector cross-hatching. | OK | Script ran without a problem. No problem detected during the quick-look. |  | Day2011053/5569-5580
22.02.2011 | 16:45 | 17:15 | GWA tilt sensor characterization (PVS #84) | n/a |  |  | n/a
22.02.2011 | 17:15 | 18:15 | Reboot of the DHAS machine | n/a | The ingestion was getting really, really slow. Next procedure started while the DHAS is still checking its disks. Exposures will be queud on the SITS in the meantime. |  | n/a
22.02.2011 | 17:32 | 1:30 | IFU-COMBO1 Full set of calibration (flat-field, wavelength…) calibration exposures for all the spectroscopic configurations and in IFU mode. | OK | Script ran without a problem. Two exposures had to be ingested manually on the DHAS (no impact on the data). No problem detected during the quick-look. |  | Day2011053/5581-5647
23.02.2011 | 1:40 | 5:10 | COMBO1 Full set of calibration (flat-field, wavelength…) calibration exposures for all the spectroscopic configurations and in MOS mode. | OK | Script ran without a problem. The selected failed-open micro-shutter for some window mode exposures turned out to be 2 micro-shutters. No other problem detected during the quick-look. We will the data as-is, we do not plan to rerun with a different selction of micro-shutter. |  | Day2011054/5648-5704
23.02.2011 | 5:21 | 13:30 | SLIT-COMBO1 Full set of calibration (flat-field, wavelength…) calibration exposures for all the spectroscopic configurations and in SLIT mode. | NOT OK | Script ran smoothly except for one exposure where the SCA491 exposure was missing (NCR raised; no explanation for that problem). For several exposures (20 out of more than 160…), there was an error in the window size and/or positioning and we will need a partial rerun. Note: partial rerun performed on 24.02.2011 |  | Day2011054/5705-5819
23.02.2011 | 13:50 | 20:02 | IMA-PSF - RUN #02 (PVS #83) Characterization of the imaging-mode PSF using the RCSS point source in the A1600 aperture. Measurements for various focus positions. | OK | Rerun of the procedure with a fainter RCSS source (MOSBB_N instead of MOS1_N). Script ran smoothly and no problem was detected during quick-look. New focus curves were generated for the RCSS confirming the best focus position derived during the preparation phase (-680 microns to be compared to the -700 microns of the preparation phase, with a tolerance of typically 100 microns). RCSS PSF exhibits some unexpected features away from the core. More analysis needed. |  | Day2011054/5820-5945
23.02.2011 | 20:07 | 1:43 | IMA-DIST Characterization of the distortion within the IFU field of view (scope limited to the IFU in cycle 1). | ONE EXPOSURE IS MISSING NO RERUN | We lost one of the first exposures (command failure; command executed gain but only the SCA492 exposure came through; one unique ID without any exposure associated). After that, the script ran very smoothly. IFU images were reconstructed and confirmed that the dither patern was OK. |  | Day2011054/5946-6024
24.02.2011 | 1:45 | 3:34 | MON-DET-DARK-SHORT #04 Sequence of 10 short dark exposures (88 groups). | OK | Script ran without a problem but once more the unzipping of one of the files on the DHAS was terribly long (final duration of the procedure is 3 hours rather than the expected 2 hours). No problem detected during quick-look. |  | Day2011054/6025 Day2011055/6026-6029
24.02.2011 | 3:37 | 4:45 | SLIT-COMBO2 Characterization of the response of the instrument for a point source (RCSS) in the A1600 aperture. Characterization of the spatial PSF in spectroscopic mode. | NOT OK | Script was running without a problem but another window size/location problem was detected. Procedure was interrupted. Note: the procedure was updated immediately and rerun (see below). |  | Day2011055/6030-6039
24.02.2011 | 4:54 | 5:20 | SLIT-COMBO2 - RUN #02 (PVS #86) Characterization of the response of the instrument for a point source (RCSS) in the A1600 aperture. Characterization of the spatial PSF in spectroscopic mode. | NOT OK | The SLIT-COMBO2 procedure was updated to increase the size of the window. However, the number of groups was not updated and this rerun was also interrupted. Note: the procedure was updated immediately and rerun successfully this time (see below). |  | Day2011055/6040-6041
24.02.2011 | 5:25 | 8:25 | SLIT-COMBO2 - RUN #03 (PVS #86) Characterization of the response of the instrument for a point source (RCSS) in the A1600 aperture. Characterization of the spatial PSF in spectroscopic mode. | OK | Procedure ran smoothly and no problem was detected during quick-look. |  | Day2011055/6042-6083
24.02.2011 | 8:31 | 8:56 | MON-DIST-SPECTRO Monitoring exposure (limited scope for cycle 1): MSA to FPA coordinate transform. | OK | Procedure ran smoothly and no problem was detected during quick-look. |  | Day2011055/6084-6085
24.02.2011 | 9:15 | 10:24 | SLIT-COMBO1-SUB (partial rerun of SLIT-COMBO1; PVS #85) Procedure created for a partial rerun of SLIT-COMBO-1. Full set of calibration (flat-field, wavelength…) calibration exposures for a subset of spectroscopic configurations and in SLIT mode. | OK | Procedure ran smoothly and the window positions/sizes are OK. |  | Day2011055/6086-6105
24.02.2011 | 10:33 | 12:07 | MON-MSA-FPA-RAD - RUN #03 Monitoring exposures (limited scope for cycle1): failed-open map of the MSA; detector cross-hatching. | OK | Procedure ran smoothly and no problem was detected during quick-look. |  | Day2011055/6106-6117
24.02.2011 | 12:16 | 14:10 | CAA-LINE3-DUTY-CYCLE (PVS #87) Procedure to test the behaviour of the CAA/LINE3 lamp when operated with a heavy duty cycle (100s ON / 200s OFF). | OK | Procedure ran smoothly. Temperature of the CAA telescopes was monitored and equilibrium was reached after one hour. Procedure was interrupted after 2h (after exposure 39). No obvious change in the lamp spectrum detected. |  | Day2011055/6118-6156

## Stabilization to HOT cond.
Date | Start time | End time | Activities / Envents | Status | Comments | CSV file(s) | Range of exposure IDs
ACTIVITIES CONDUCTED BETWEEN THE CALIBRATION AND CONFIRMATION PHASES (STABILIZATION TO HOT CONDITIONS)
25.02.2011 | 1:00 | 15:45 | BENCH WARM-UP PHASE As the bench was warming up, series of 3 exposures were obtained with NIRSpec in the OPAQUE/MIRROR configuration: an OFF exposure , a FLAT5-source exposure (bright source) and a TEST-source exposure (regular imaging). Monitoring of these expores will allow us to look for displacement of the image of the MSA on the FPA as the temperature changes. CAUTION: during the initial part of the bench warm-up the detector got warmer also and detector temperature control was lost during some time. | n/a | No quick look performed. |  | Day2011056/6166-6213
25.02.2011 | 18:00 | 19:10 | DS FUNCTIONAL TESTS Activity to be confirmed. | n/a | No quick look performed on ESA side. |  | Day2011056/6214-6226
25.02.2011 | 22:15 | 8:15 | BENCH WARM-UP PHASE As the bench was warming up, series of 3 exposures were obtained with NIRSpec in the OPAQUE/MIRROR configuration: an OFF exposure , a FLAT5-source exposure (bright source) and a TEST-source exposure (regular imaging). Monitoring of these expores will allow us to look for displacement of the image of the MSA on the FPA as the temperature changes. | n/a | No quick look performed. |  | Day2011056/6227-6239 Day2011057/6240-6262
28.02.2011 | 21:30 | 7:32 | FILL-DET-DARK-SHORT - RUN #01 (PVS #96) Long sequence of 30 "short" (88 groups) dark exposures. | OK | Script ran smoothly. As usual there was some backlog at the end but nothing unreasonable. Nothing detected at quick-look. |  | Day2011059/6263-6274 Day2011060/6275-6292
01.03.2011 | 7:39 | 8:18 | DET-WINDOW-TEST-SUB (PVS #96) Comparison of the performances of the detector when using the full-frame / window / stripe modes (hopefully, no difference). Partial rerun of some exposures that failed during the calibration phase. | OK | This time we got window-mode exposures and not full-frame ones. There were timeout errors for each exposures (NCR-1199) due to an erroneous estimation of the integration time for this type of window. This had no impact on the data. |  | Day2011060/6293-6295

## Confirmation phase
Date | Start time | End time | Activities / Envents | Status | Comments | CSV file(s) | Range of exposure IDs
ACTIVITIES CONDUCTED DURING THE CONFIRMATION PHASE
01.03.2011 | 8:32 | 10:08 | MON-MSA-FPA-RAD - RUN #01 (PVS #96) Monitoring exposures (limited scope for cycle1): failed-open map of the MSA; detector cross-hatching. | OK | Script ran smoothly and no problem was detected during quick-look (except when the printer ran out of toner mimicking nice stripes on the detectors…). IMPORTANT: when we conducted this test, the CMO database included a 100 micron correction along the z-axis (focus) with respect to what we used during the calibration phase. This should have no detectable effect but it is there. |  | Day2011060/6296-6307
01.03.2011 | 10:10 | 11:00 | Restoring the database to its calibration phase state | n/a | Starting at this point the CMO coordinate system is identical to the one used during the calibration phase. |  | n/a
01.03.2011 | 11:00 | 14:12 | IMA-FOCUS-SHORT Determination of the best focus RMA positions for the nominal OTE focal plane position. Short version of the procedure meant to provide exposures for the comparison with the IMA-FOCUS procedure of the calibration phase. | OK | Script ran smoothly and the pinholes were there. Preliminary analysis indicates a nominal focus position around -380 microns but with quite a large uncertainty (of the order of 100 microns). Difficult to say if this is really different from the calibration phase value of -280 microns as we are really on the flat part of the curve. |  | Day2011060/6308-6337
01.03.2011 | 14:17 | 14:54 | MON-DIST-SPECTRO Monitoring exposure (limited scope for cycle 1): MSA to FPA coordinate transform. | OK | Script ran smoothly and no problem is detected during quick-look. |  | Day2011060/6338-6339
01.03.2011 | 15:00 | 15:31 | BIG-WINDOW-CHECK - RUN #01 - PVS#100 | NOT OK | Trouble-shooting for the correction of the bug detected while executing DET-WINDOW-TEST-SUB (NCR-1199). Problem is still there (the fix was actually not implemented). Rerun needed once the fix is implemented. |  | Day2011060/6340-6342
01.03.2011 | 15:36 | 16:06 | BIG-WINDOW-CHECK - RUN #02 - PVS#100 | OK | Trouble-shooting for the correction of the bug detected while executing DET-WINDOW-TEST-SUB (NCR-1199). Problem has been fixed. |  | Day2011060/6343-6345
01.03.2011 | 16:09 | 0:55 | GWA-SENSOR-CALIB (modified as per PVS#102) Characterization of the repeatability of the GWA for its different positions. Calibration of the GWA tilt sensors. | OK | Script ran without problem. No problem detected during quick-look. Preliminary analysis shows a good correlation between shifts and tilt sensor output. |  | Day2011060/6346-6495 TBD (may be missing in the archive)
02.03.2011 | 0:49 | 2:38 | IMA-DIST-SHORT Characterization of the distortion within the IFU field of view (scope limited to the IFU in cycle 1). Short version of the procedure meant to provide exposures for the comparison with the IMA-DIST procedure of the calibration phase. | OK | Script ran without problem. Dither pattern is OK. Maybe a very small shift of the position of the pinhole within the IFU field of view (order of 1 slice). |  | TBD (missing in the archive)
02.03.2011 | 2:42 | 5:47 | MON-DET-DARK-SHORT RUN #01 Sequence of 10 short dark exposures (88 groups). | OK | Script ran without problem but the ingestion was slow like hell. |  | TBD (missing in the archive)
02.03.2011 | 5:49 | 18:33 | IFU-COMBO1 (replaces MON-COMBO-IFU as per PVS #101) Full set of calibration (flat-field, wavelength…) calibration exposures for all the spectroscopic configurations and in IFU mode. | OK | Script ran without a problem except for the failure of one exposure. It was possible to repeat it. No issue detected during quick-look. Light level similar to those of the calibration phase. |  | Day2011061/6523-6584 TBD (missing in the archive)
02.03.2011 | 18:45 | 22:21 | COMBO1-01 (replaces MON-COMBO-MOS as per PVS #101) Full set of calibration (flat-field, wavelength…) calibration exposures for all the spectroscopic configurations and in MOS mode. | OK | Script ran without a problem. Window sizes and location are OK. |  | Day2011061/6673-6729
02.03.2011 | 22:23 | 22:53 | STRAYLIGHT-ARC (PVS #105) Special set of exposure aiming at characterizing the straylight arc present in the high-spectral resolution configurations. In particular, this test sequence includes "straylight-only" exposures. | OK | Script ran without a problem. Very interesting images. |  | TBD (missing in the archive)
02.03.2011 | 23:02 | 0:42 | MON-MSA-FPA-RAD - RUN #02 Monitoring exposures (limited scope for cycle1): failed-open map of the MSA; detector cross-hatching. | OK | Script ran smoothly and no problem was detected during quick-look. |  | TBD (missing in the archive)
03.03.2011 | 0:43 | 2:44 | RCSS-HALO (PVS #106) Ad-hoc test sequence aiming at characterizing the halo observed around the RCSS SPF. | OK | Script ran smoothly and no problem was detected during quick-look. A lot of questions, not a lot of answers. |  | Day2011061/6749-6769